## Slide 1
1. Institute of Micro- and Nanostructure Research, Friedrich-Alexander-Universität Erlangen-Nürnberg
2. Department of Chemistry, Indiana University Bloomington
Automated Tomographic 4D-STEM Data Acquisition Workflow Connected with Ptychographic HR 3D Reconstruction
Shengbo You¹, Baixu Zhu², Xingchen Ye², Philipp Pelz¹

## Slide 2
Overview
Motivation
Configuration
Workflow
Precision analysis
Demonstration with
Reconstruction analysis

## Slide 3
Multi-slice ptychography achieves very high 2d resolution, but limited axial resolution
- 3D resolution can be enhanced by tilt series experiments
- We have developed Ptychographic tomography as a powerful method to achieve 3D atimc resolution phase-contrast imaging
But: requires highly trained microscopist & long experiments

## Slide 4
End to end reconstruction workflow
You et al., Phys. Scr., 2024, 100, 015404.
Recover the 3D structure of the sample directly from the 4D STEM data

## Slide 5
Motivation
Dose efficiency: nearly all electrons that pass through the target area are contributing as data.
Hands-free expertise: achieve atomic resolution level data acquisition without well trained operator.
Real time feedback: check the data quality immediately after acquisition.
Our goal: automated tomographic 4D STEM data acquisition achieving atomic resolution
Aberration
Phase projection
Multislice
Carbon accumulation

## Slide 6
SerialEM
Parallax
4D STEM compressed data
Scan API
Image/data
request
4D-STEM
Camera Plugin
Image/data
request
Multislice ptychography
Defocus
Custom Scripts
SerialEM
command
Synchronised scan signal & trigger
Microscope PC
High performance
Workstation
End to end
reconstruction
Alignment
2D phase projection
Affine
transformation
HAADF image
Image/data request
HAADF
image
TEM scripting
interface
4D STEM raw data
Detector control PC
Arina Detector
Sample
Scan coil
HAADF
detector
Scan Controller
HAADF signal
HAADF image
Custom script control TEM to perform the data acquisition. Data is reconstructed in real time

## Slide 7
Workflow
Ground truth at log mag (57k)
Ground truth of reference region
Record the data
Tilt to the next angle
Course alignment at low mag (57k)
Take image at reference region
Calculate the defocus and offset
Initial degree
Fine alignment at reference area
Two level of alignment to ensure fine tracking of the target area

## Slide 8
Workflow demonstration
Course alignment (low mag)
Fine alignment (reference area)
data acquisition
5nm
Nearly all electrons went through the target area are contributing as data

## Slide 9
High-precision automated tracking
Data collection at target area
± 10 nm tracking error through the tilt series

## Slide 10
Sample:
2nm
5nm
Liu et al., J. Am. Chem. Soc., 2020, 142, 12777–12783.
Gd₂O₃ nanohelices formed via bilayer lattice misfit strain.

## Slide 11
Multislice and linear tomography
Projection of each angle
Multislice ptychography reconstruction show atomic resolution in 2D

## Slide 12
End to end reconstruction
1nm³
Reconstructed 3D volume
Cross-section
End to end reconstruction show atomic resolution in 3D

## Slide 13
Result analysis
1.05Å
0.88Å
0.81Å
Reconstructed volume in Fourier domain
Sub- Ångstrom 3D resolution

## Slide 14
conclusion
Tomographic 4D STEM with high dose efficiency
Atomic resolution data acquisition without well-trained operator
Less than 10nm error
Sub-(Å)ngstrom 3D resolution

## Slide 15
Acknowledge
Indiana University Bloomington
Sample preparation
Supervision
Friedrich-Alexander-Universität Erlangen-Nürnberg
Baixu Zhu
Prof. Xingchen Ye
Prof. Dr. Philipp Pelz
Hardware support
Matthias Meffert